## Standard Telegrams to send-rcv
Category |  | IONICS FW Name | Turbo Pump Command | Leybold | [11 empty] | Comments | Values always returned: |  |  |  |  | Warning Bits:
|  |  |  | 0 1 | 2 3 | 4 5 | 6 7 | 8 9 | 10 11 | 12 13 | 14 15 | 16 17 | 18 19 | 20 21 | 22 23
?? |  |  | Bearing Break-in function (P119) | 216 | 10 | 7700 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 76 | Not sure what this is | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 0 | Pump temp1 surpassed | Parameter 312: |  | HEX | ASCII
configuration | 1.2 | Pump Model | Product Name (Index ASCII), P313 | 216 | 61 | 3900 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 49 | Must request Field value for indexed values. Must request index values from 0-17. Convert to ASCII to read string. See Example. | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | TMP 350/450 iC | 3 | mimimum ambient Temp not reached |  | 2 | 30 | 0
configuration | 1.5 | Pump SN | Pump Serial Number (P355) | 216 | 61 | 6300 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 13 | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 5 | Not Used |  | 4 | 30 | 0
configuration | 4 | Controller Model | Converter HW Version (P316) | 216 | 61 | 3C00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 004C | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 11 | Overload Warning |  | 10 | 38 | 8
configuration | 5 | Controller SN | Converter SN, (P315) | 216 | 61 | 3B00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 004B | ?? For index number, convert to ASCII 0-10 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 10 | Not Used |  | 9 | 33 | 3
configuration | 6 | Comm. Electronics Model | Communication Electronics HW Ver (P396) | 216 | 61 | 8C00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00FC | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 12 | pump temp 5 surpassed |  | 11 | 0 | #VALUE!
configuration | 6.3 | Comm. Electronics SN | Communication Electronics SN (P395) | 216 | 61 | 8B00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00FB | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  |  |  |  | 14 |  | #VALUE!
configuration | 6.5 | SW Version | Drive Electronics SW Ver (P1100) | 216 | 14 | 4C00 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 49 | Numeric Value | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 14 | PS voltage warning |  | 13 |  | #VALUE!
configuration | 10 | Pump Parameter Set | Pump Parameter Set (P349) | 216 | 61 | 5D00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 002B | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  |  |  |  | 17 |  | #VALUE!
configuration | 11 | Not Used | Read Hardware Version/Device Type | 216 | 10 | 100 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 0 | Reads Back B4 in Byte 10, =180=Turbovac350/450i | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 1 | pump temp 2 surpassed | Index | 0 | 33 | 3
configuration | 12 | Not Used | Configuration Text, Index,ASCII), P314 | 216 | 61 | 3A00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 004A | ?? For index number, convert to ASCII 0-26 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 4 | Not Used |  | 3 | 30 | 0
configuration | 13 | Not Used | Pump Serial Number (P1035) | 216 | 64 | 0B00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 75 | ?? For index, convert ASCII 0-17, same as P355? | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 6 | Overspeed warning >10Hz over |  | 5 | 39 | 9
configuration | 14 | Not Used | Pump Part Number (P350) | 216 | 61 | 5 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 002E | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 7 | Pump temp 4 surpassed |  | 6 | 39 | 9
configuration | 15 | Not Used | Pump Type/Rotor Type (P23) | 216 | 10 | 1700 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 16 | 0=TURBOVAC i/Ix CL (Classic) 1=TURBOVAC i/iX WR (wide-range) 2=TURBOVAC i/iX MI (multi-inlet) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  | 8 | Not Used |  | 7 | 30 | 0
configuration | 16 | Not Used | Converter PN (Index ASCII), P312 | 216 | 61 | 3800 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 48 | Must request Field value for indexed values. Must request index values from 0-17. Convert to ASCII to read string. See Example. | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | 6416338 | 9 | Not Used |  | 8 | 30 | 0
configuration | 17 | Not Used | Communication Electronics PN (P394) | 216 | 61 | 8A00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00FA | ?? For index, convert ASCII 0-17 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  |  |  |  | 15 |  | #VALUE!
Error |  |  | Clear Error | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage |  |  |  |  | 23 |  | #VALUE!
Error |  |  | Read Error Code | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Error Counter Total (P40) | 216 | 10 | 2800 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 29 | Counts error messages | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Overload Counter (P41) | 216 | 10 | 2900 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 28 | Counts Overload Errors | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Power Failure Counter (P43) | 216 | 10 | 2B00 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 002A | Counts Power Failure Errors | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Error Code Memory (P171) | 216 | 10 | AB00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00AA | Add index value of 0 to 253 to get individual error codes up to 254 error codes. Can use P40 (error counter) to determine how many errors to list when looking at pump history. ??=hex for 0-253. | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Rotational Frequency at time of the error (P174) | 216 | 10 | AE00 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00AF | Access analagously to P171 | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Operational Hours at the time of the error (P176) | 216 | 10 | B0000 | ??00 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00B1 | Access analagously to P171, units = 0.01Hours | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Error |  |  | Active Warnings (bit by bit) (P227) | 216 | 10 | E300 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 0 | Read Binary from Bytes 7-10, compare with table | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status | 1 | Decelerating Flag | Decelerating | Leybold 0 1: Always read with any command
Status | 2 | Accelerating flag | Accelerating | Leybold 0 1: Always read with any command
Status | 3 | Normal Operation Flag | Normal Operation | Leybold 0 1: Always read with any command |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status | 4 | Status Bar with Speed | Read Actual Rotation Speed in Hz | Leybold 0 1: Always read with any command |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status | 5 | Normal Operation Flag plus speed at 1000 Hz | Read Set Rotation Speed Attained | Leybold 0 1: With any command, bit 10 enabled in Bytes 11-12) |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status | 6 | Normal Operation Flag | Read Rotation Switch Point Attained | Leybold 0 1: With any command, bit 2 enabled, bits 4-5 disabled in Bytes 11-12) |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status | 7 | Number of Starts | Number of Starts (P38) | 216 | 10 | 2600 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 27 | Number of times the pump has been started | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Level (1000Hz) used in FW logic, but parameter itself is not used (preset in pump) | Nominal Rotational Frequency (P18) | 216 | 10 | 1200 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 13 | (1000 Hz) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Not Used | Minimum Nominal Frequency (P19) | 216 | 10 | 1300 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 12 | (500 Hz) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Level monitored during startup, but parameter not used | Minimum Frequency Level (P20) | 216 | 10 | 1400 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 15 | (490 Hz) must pass this point within 500 s after start initiated. | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Time monitored during startup, but parameter is not used. | Max. Passing Time (P183) | 216 | 10 | B700 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00B6 | Maximum time for minimum frequency level to be reached (500 s) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Not Used | Setpoint Frequency (P24) | 216 | 10 | 1800 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 19 | (1000 Hz) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Not Used (only for changing setpoint, 90% is fine) | Normal Operation %Setpoint (P25) | 216 | 10 | 1900 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 18 | % of setpoint where normal operation can be designated above (90%) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
Status |  | Time monitored during startup, but will not change, not use parameter | Max Run Up Time (P32) | 216 | 10 | 2000 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 21 | Max time to get to (90%) rotation (in seconds)--2000s (33min,20s) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage | [6 empty] | #VALUE! | #VALUE!
General |  |  | Standby Frequency (P150,800 Hz) | 216 | 10 | 9600 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 97 | Not sure when and why we would use this | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Power | 1 |  | Read Motor Voltage in V | Leybold 0 1: Always read with any command |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Power | 2 |  | Read Motor Power in W | 216 | 10 | 600 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 7 | Reads units of 0.1 response = 1256=125.6W | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Power | 3 |  | Read Motor Current in A | Leybold 0 1: Always read with any command |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
power | 4 |  | Mean Pump Current (P140) | 216 | 10 | 8C00 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 008D | Average Current Consumption (0.1 A) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
power | 5 |  | Maximum Current (P17) | 216 | 10 | 1100 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 10 | Reads back max allowable current in 0.1A units | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Power |  |  | Motor Current threshold (P21) | 216 | 10 | 1500 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 14 | % of current, if exceeded after normal operation attained, produces "High Load Error" | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
power |  |  | Max converter DC Input Current (P185) | 216 | 10 | B900 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00B8 | 9.0 A (0.1A) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Start/Stop |  |  | Turn Pump Motor On | 216 | 0 | 0 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 11 | Starts the pump - need to ensure used for polling every 3 seconds to keep pump turning. | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Start/Stop |  |  | Turn Pump Motor Off | 216 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 10 | Stops the pump - need to ensure the start byte is not resubmitted, restarting the pump until the power is disabled to the turbo. | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 1 | Bearing Temperature | Read Bearing Temperature (P125) | 216 | 10 | 7D00 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 007C | Reads Bearing Temp (degC) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 2 | Motor Temperature | Read Motor Temperature | 216 | 10 | 700 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 6 | Reads back in degrees C | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 3 | Controller Temperature | Actual Frequency Converter Temp. | Leybold 0 1: Always read with any command |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 4 | Limit Used, but Parameter Not Used | Motor Lower Temperature Error Level (P131) | 216 | 10 | 8300 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 82 | Lower Motor temperature error limit (0C) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 5 | Limit Used, but Parameter Not Used | Motor Lower Temperature Warning Level (P128) | 216 | 10 | 8000 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 81 | Lower motor temperature warning Limit (5C) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 7 | Limit Used, but Parameter Not Used | Motor Upper Temp Warning Level (P16) | 216 | 10 | 1000 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 11 | Upper Warning Limit for Motor Temperature (80C) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 8 | Limit Used, but Parameter Not Used | Upper Motor Temperature Error Level (P133) | 216 | 10 | 8500 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 84 | Upper motor temperature error limit (100C) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 10 | Limit Used, but Parameter Not Used | Freq. Converter Temp error limit (P1102) | 216 | 14 | 4 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 004B | degC | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 11 | Duplicate | Freq. Converter Temp Warning Level (P1101) | 216 | 14 | 4D00 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 48 | degC | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 12 | Limit Used, but Parameter Not Used | Bearing Temp Warning Level(P126) | 216 | 10 | 7 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 007F | Warning Limit for Bearing Temperature (60C) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Temperature | 13 | Limit Used, but Parameter Not Used | Bearing Temp Error Level (P132) | 216 | 10 | 8400 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 85 | Error Limit for Bearing Temperature (65C) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Status |  |  | Read Operating Hours in h | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
Status |  |  | Pump Controller Operating Hours (P184) | 216 | 10 | B800 | 0 | 0 | 4 | 100 | 0 | 0 | 0 | 0 | 00B9 | Counts the hours on the controller (0.01 Hours) | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216; Leybold 22 23: z |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
|  |  |  | Leybold 0 1: 216 |  | Frequency | Freq. Converter Temp | Motor Current | PS Voltage
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216
216