## Slide 1
Initializing
Plug system into the wall
Mass Spectrometer is Off
Turn power on (from the software or the button)
Analyzer Off
Press Start
Starting Up
Set all DAC channels to 0, except for curtain gas
Starting Analyzer
Load oscillators
Check card cage door
(A120 > 0)
Starting Analyzer
Card Cage Door Open
Check curtain gas
(A64 > 0.35V)
Starting Analyzer
Check Drying Gas
Check Ion Source Interlock
Card cage door open. Stay at this stage until door is closed
No curtain gas. Stay at this stage until curtain gas is enabled
Continue to Turbo Pump Ready state but not setting HSID temperature.
Turn on roughing pump and set HSID temp, wait up to 30 seconds
Check foreline Pressure (A74 < 0.35V)
Foreline pressure failure, go to Mass Spectrometer is Off state on the 3rd failure .
Starting Analyzer
Foreline Press. High
Roughing Pump Ready
Turning Power Off
Starting Analyzer
Startup Cancelled
Turbopump Ready
Turbo Startup
Starting Analyzer
(Pumping 1 progress bar)
Turn on turbo pump, wait up to 20 minutes to reach 80% speed
Check turbo speed
(RS >690 Hz)
Turbo not up to speed, go to Roughing Pump Ready state.
Starting Analyzer
Vacuum Failure 002
Check vacuum gauge
(90 < A119 < 153)
Overpressure . System Shutdown, go to Roughing Pump Ready state.
Starting Analyzer
Vacuum Failure 016
Turning Turbo Off
Checking Vacuum
Starting Analyzer
Startup Cancelled
Starting Analyzer
(Pumping 2 progress bar)
Check vacuum gauge and vacuum ready
(90 < A119 < 145)
(A121 = 0)
Check Instrument
Vacuum Failure 003
High vacuum not achieved after 30 minutes, go to Turbopump Ready state.
Check Ion Source Interlock
Enable HV and set all DAC channels to default values
Going to Ready
Analyzer Ready
Ready
System Polling
(check all interlocks)
Source Standby
Possible LCD Messages
Startup Sequence
No
No
No
No
No
No
No
No
Yes
Yes
Yes
Yes
Yes
Yes
Yes
Yes
(Repeat 2 times if fails)