## DSV
Auto Cal-Res Utility |  | Status
ITEM | RESULT | V 4.8.6 | COMMENTS/SOLUTION
USER INTERFACE AND USABILITY
Abort option needed | Not in current version | NEW | Right now there is no way to abort for any reason. If the AutoCal box is closed using the close gadget Ionica Server hangs so badly Task manager is needed.
Pause/Continue option needed | Not in current version | NEW | What if we need to refill a syringe?
Peaks present validated | Not in current version | BUG | All peaks selected must be present, If not, abort with informative error message.
Individual RT Plots Correctly labelled | No | BUG | Why do the titles read Mass 175.13, 0?
Individual RT Plots Correctly labelled | No | BUG | For unused windows please leave title bar blank
Start button to initiate Calibration Routine | No | BUG | Final layout to be specified in Ionica Server Redesign specs but button should be labelled "Start", not "Calibration".
Validation
Reminder to use ESI Side 1 or as appropriate for saved default method | Not in current version | NEW | I forgot and used side 2 for my tests! Proper peak validation should deal with this adequately.
MassCalibrationReference.cal file validated | Not in current version | BUG | File was present but list of masses not correct for Series. Utility gave error box then crashed/hung. Must validate for both presence of file, and content/format. In case of problem, abort with informative error message.
Only selected peaks used
Peaks validated | No | BUG
Q1 Calibration
Terminates and proceeds to Q2 | No | BUG | Server hangs. Error box generated. Task manager needed to abort.
Q2 Calibration
xx
Results Table
Reference masses listed correctly | No | BUG | Mass initally displayed reads 59.06. Table mass is 59.05.
Report
Search Parameters reported correctly | Yes | OK
Number of Cycles reported/used correctly

## Package
Dual Source Validation Package |  | Status
ITEM | RESULT - Current Version | Version | COMMENTS/SOLUTION
REPORTING (Shahram) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - Real Time DataViewer (Alex M) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - DataViewer (George) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - Data Post-Processing (Stephen O'H) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - Molana Data Acquisition (PDI - Defer) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly