## DSV
Auto Cal-Res Utility |  | Status
ITEM | RESULT | V 4.8.6 | COMMENTS/SOLUTION
USER INTERFACE AND USABILITY
Abort option needed | Not in current version | NEW | Right now there is no way to abort for any reason. If the AutoCal box is closed using the close gadget Ionica Server hangs so badly Task manager is needed. | Done
Pause/Continue option needed | Not in current version | NEW | What if we need to refill a syringe? | Done
Peaks present validated | Not in current version | BUG | All peaks selected must be present, If not, abort with informative error message. | ?
Individual RT Plots Correctly labelled | No | BUG | Why do the titles read Mass 175.13, 0? | Done
Individual RT Plots Correctly labelled | No | BUG | For unused windows please leave title bar blank | Done
Start button to initiate Calibration Routine | No | BUG | Final layout to be specified in Ionica Server Redesign specs but button should be labelled "Start", not "Calibration". | Done
Validation
Reminder to use ESI Side 1 or as appropriate for saved default method | Not in current version | NEW | I forgot and used side 2 for my tests! Proper peak validation should deal with this adequately. | ?
MassCalibrationReference.cal file validated | Not in current version | BUG | File was present but list of masses not correct for Series. Utility gave error box then crashed/hung. Must validate for both presence of file, and content/format. In case of problem, abort with informative error message. | ?
Only selected peaks used
Peaks validated | No | BUG
Q1 Calibration
Terminates and proceeds to Q2 | No | BUG | Server hangs. Error box generated. Task manager needed to abort. | Done
Q2 Calibration
xx
Results Table
Reference masses listed correctly | No | BUG | Mass initally displayed reads 59.06. Table mass is 59.05. | ?
Report
Search Parameters reported correctly | Yes | OK
Number of Cycles reported/used correctly

## Package
Dual Source Validation Package |  | Status
ITEM | RESULT - Current Version | Version | COMMENTS/SOLUTION
REPORTING (Shahram) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - Real Time DataViewer (Alex M) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - DataViewer (George) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - Data Post-Processing (Stephen O'H) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly
DISPLAY - Molana Data Acquisition (PDI - Defer) |  | NEW
Source Type correct (ESI, APCI)
Source ID correct (Side 1 or Side2)
Source Polarity correct (Positive vs. Negative)
Source Voltage correct (ISV, CD)
Delay times reported/used correctly
Number of Cycles reported/used correctly